## Test Overview
Tip Force Testing | testing the ccm notched wrist vs. rectangular notched wrist
Test | Description | Purpose
Blocking Force | attached cable distal to the joint and pulled on cable perpendicular to the notch using the force sensor to record how much force is pulling distal to the notch perpendicularly and thus how much the tip is being displaced as a result of that. | when handling tissues within the middle ear during surgery, is the tip able to handle those forces need different directions
Cyclic Loading |  | to determine how many times the joint can be loaded and unloaded before it starts to fail/fracture/crack
Twisting |  | degrees the joint can be twisted before it cracks

## Blocking Force
| Blocking Force vs. Cable Displacement |  |  | Notes: - use fresh notch, not plastically deformed - 0.3mm cable - measure distance between clamp and notch - measure distance between top of notch and where force is applied to the tip
Sample | Blocking Force (N) | Tip Displacement (mm) | Cable Displacement (mm)
CCM v-1 2 notches |  |  | 0
distance between clamp and notch: 4.4mm | 0.004 | 0 | 0.025
distance between top of notch and point force is applied to the tip (position of the hole) | 0.006 | -0.001 | 0.05
laser hitting approximately on opposite side of tube where the hole is | 0.008 | 0.001 | 0.075
| 0.01 | 0.002 | 0.1
| 0.0126 | 0.004
| 0.0144 | 0.01
| 0.0167 | 0.009
| 0.0187 | 0.003
| 0.02 | 0.004
| 0.023 | 0.008
| 0.025 | 0.01
| 0.028 | 0.01
| 0.032 | 0.01
| 0.036 | 0.012
| 0.037 | 0.013
| 0.04 | 0.007
| 0.044 | 0.012
| 0.047 | 0.01
| 0.051 | 0.013
| 0.054 | 0.016
| 0.058 | 0.02
| 0.061 | 0.024
| 0.066 | 0.03
| 0.07 | 0.038
| 0.079 | 0.048
| 0.084 | 0.05
| 0.088 | 0.053
| 0.086 | 0.053
| 0.094 | 0.06
| 0.098 | 0.06
| 0.103 | 0.07
| 0.108 | 0.075
| 0.113 | 0.08
| 0.119 | 0.085
| 0.125 | 0.1
| 0.133 | 0.105
| 0.14 | 0.11
| 0.15 | 0.12
| 0.158 | 0.125
| 0.163 | 0.12
| 0.174 | 0.13
| 0.185 | 0.14
| 0.195 | 0.145
| 0.2 | 0.16
| 0.216 | 0.18
| 0.232 | 0.22
| 0.24 | 0.22
| 0.25 | 0.23
| 0.26 | 0.24
| 0.27 | 0.27
| 0.283 | 0.29
repositioned to ensure the joint is bending in the horizontal plane | 0.24 | 0.3
| 0.258 | 0.32
| 0.27 | 0.33
| 0.286 | 0.35
| 0.3 | 0.36
| 0.312 | 0.38
| 0.316 | 0.38
| 0.34 | 0.4
| 0.36 | 0.44
| 0.38 | 0.47
| 0.41 | 0.49
| 0.41 | 0.49
| 0.43 | 0.51
| 0.45 | 0.53
| 0.475 | 0.58
| 0.497 | 0.63
| 0.52 | 0.7
| 0.547 | 0.75
| 0.58 | 0.8
| 0.6 | 0.85
| 0.62 | 0.9
| 0.65 | 0.96
| 0.68 | 1.04
| 0.7 | 1.07
| 0.73 | 1.16
| 0.755 | 1.23
| 0.782 | 1.3
| 0.81 | 1.43
| 0.84 | 1.53
| 0.86 | 1.6
| 0.89 | 1.74
| 0.91 | 1.85
| 0.925 | 1.93
| 0.97 | 2.05
| 0.98 | 2.2
| 1 | 2.4
| 1.035 | 2.6
| 1.06 | 2.7
| 1.09 | 2.9
| 1.13 | 3.05
| 1.15 | 3.3
| 1.17 | 3.4
| 1.16 | 3.5
| 1.23 | 3.6
| 1.27 | 3.8
| 1.29 | 3.98
| 1.36 | 4.1
| 1.41 | 4.2
| 1.33 | 4.4
| 1.19 | 4.1
| 1.095 | 3.9
| 1.01 | 3.6
| 0.95 | 3.4
| 0.9 | 3.1
| 0.85 | 3
| 0.81 | 2.75
| 0.78 | 2.5
| 0.74 | 2.35
| 0.705 | 2.15
| 0.68 | 2.03
| 0.67 | 1.9
| 0.644 | 1.8
| 0.61 | 1.7
| 0.6 | 1.6
| 0.55 | 1.45
| 0.536 | 1.3
| 0.515 | 1.3
error in laser noticed | 0.497 | 1.15
1
fixed setup | 0.01 | 0.001
distance between clamp and notch: 4.4mm | 0.02 | 0.01
distance between top of notch and point force is applied to the tip (position of the hole) | 0.03 | 0.03
laser hitting approximately on opposite side of tube at tip | 0.04 | 0.04
| 0.05 | 0.073
| 0.06 | 0.077
| 0.07 | 0.086
| 0.081 | 0.096
| 0.09 | 0.14
| 0.1 | 0.177
| 0.15 | 0.267
| 0.125 | 0.23
| 0.152 | 0.28
| 0.176 | 0.318
| 0.2 | 0.37
| 0.227 | 0.4
| 0.251 | 0.45
| 0.277 | 0.54
| 0.3 | 0.62
| 0.325 | 0.657
| 0.349 | 0.73
| 0.376 | 0.753
| 0.396 | 0.75
| 0.424 | 0.81
| 0.45 | 0.87
| 0.474 | 0.91
| 0.5 | 0.95
| 0.523 | 1.045
| 0.55 | 1.11
| 0.577 | 1.17
| 0.6 | 1.22
| 0.625 | 1.26
| 0.647 | 1.34
re-positioned clamp | 0.01 | 0.0127
| 0.02 | 0.019
| 0.03 | 0.04
| 0.04 | 0.07
| 0.05 | 0.0764
| 0.06 | 0.075
| 0.0025 | 0.01
| 0.01 | 0.013
| 0.02 | 0.027
| 0.03 | 0.06
| 0.04 | 0.077
| 0.05 | 0.094
| 0.06 | 0.121
| 0.07 | 0.19
| 0.08 | 0.22
| 0.09 | 0.25
| 0.1 | 0.25
| 0.125 | 0.237
| 0.15 | 0.288
| 0.175 | 0.36
| 0.2 | 0.43
| 0.225 | 0.49
| 0.25 | 0.546
| 0.275 | 0.596
| 0.3 | 0.62
| 0.326 | 0.659
| 0.35 | 0.7
| 0.376 | 0.792
| 0.408 | 0.855
| 0.425 | 0.898
| 0.451 | 0.982
| 0.474 | 1.019
| 0.503 | 1.067
| 0.525 | 1.1
| 0.55 | 1.2
| 0.57 | 1.28
| 0.6 | 1.34
| 0.65 | 1.43
| 0.7 | 1.56
| 0.75 | 1.72
| 0.8 | 1.84
| 0.854 | 2.05
| 0.9 | 2.305
| 0.95 | 2.45
| 1 | 2.56
| 1.05 | 2.72
| 1.1 | 2.87
| 1.15 | 2.99
| 1.05 | 3.143
| 1 | 3.095
| 0.953 | 3
| 0.9 | 2.88
| 0.857 | 2.77
| 0.8 | 2.67
| 0.753 | 2.55
| 0.71 | 2.42
| 0.65 | 2.23
| 0.6 | 2.05
| 0.555 | 1.82
| 0.5 | 1.61
| 0.452 | 1.24
| 0.404 | 1.12
| 0.35 | 0.95
| 0.3 | 0.817
| 0.275 | 0.71
| 0.2 | 0.54
| 0.15 | 0.4
| 0.1 | 0.23
| 0.09 | 0.18
| 0.08 | 0.16
| 0.07 | 0.143
| 0.06 | 0.106
| 0.05 | 0.08
| 0.04 | 0.051
| 0.03 | 0.054
| 0.02 | 0.055
2 | 0.01 | 0.042
| 0.02 | 0.055
| 0.03 | 0.06
| 0.04 | 0.102
| 0.05 | 0.126
| 0.06 | 0.153
| 0.07 | 0.18
| 0.08 | 0.21
| 0.09 | 0.193
| 0.1 | 0.191
| 0.126 | 0.23
| 0.15 | 0.276
| 0.175 | 0.36
| 0.2 | 0.425
| 0.226 | 0.49
| 0.25 | 0.555
| 0.275 | 0.6
| 0.3 | 0.627
| 0.375 | 0.678
| 0.4 | 0.857
| 0.46 | 0.986
| 0.51 | 1.1
| 0.557 | 1.22
| 0.6 | 1.26
| 0.66 | 1.48
| 0.7 | 1.65
| 0.76 | 1.77
| 0.8 | 1.95
| 0.84 | 2.13
| 0.9 | 2.28
| 0.96 | 2.45
| 1 | 2.59
| 1.1 | 2.85
| 1.2 | 3.13
| 1.32 | 3.54
| 1.46 | 3.8
| 1.55 | 3.95
| 1.67 | 4.19
| 1.52 | 4.02
| 1.42 | 3.93
| 1.31 | 3.8
| 1.23 | 3.68
| 1.1 | 3.5
| 1 | 3.34
| 0.91 | 3.12
| 0.8 | 2.84
| 0.7 | 2.54
| 0.61 | 2.15
| 0.5 | 1.67
| 0.4 | 1.19
| 0.3 | 0.815
| 0.206 | 0.56
| 0.098 | 0.27
| 0.093 | 0.24
| 0.08 | 0.24
| 0.07 | 0.256
| 0.06 | 0.25
| 0.049 | 0.21
| 0.04 | 0.16
| 0.03 | 0.12
| 0.02 | 0.064
3 | 0.01 | -0.008
re-positioned tape for laser | 0.01 | -0.001
| 0.02 | 0.002
| 0.03 | 0.025
| 0.04 | 0.043
| 0.05 | 0.063
| 0.061 | 0.08
| 0.07 | 0.092
| 0.08 | 0.097
| 0.09 | 0.12
| 0.1 | 0.17
| 0.126 | 0.19
| 0.15 | 0.21
| 0.175 | 0.25
| 0.2 | 0.306
| 0.225 | 0.353
| 0.25 | 0.42
| 0.275 | 0.445
| 0.3 | 0.528
| 0.326 | 0.565
| 0.35 | 0.597
| 0.375 | 0.68
| 0.4 | 0.779
| 0.45 | 0.89
| 0.5 | 1.04
| 0.6 | 1.25
| 0.7 | 1.56
| 0.79 | 1.81
| 0.9 | 2.11
| 1 | 2.46
| 1.1 | 2.74
| 1.2 | 3
| 1.3 | 3.37
| 1.42 | 3.53
| 1.52 | 3.76
| 1.6 | 3.91
| 1.49 | 3.78
| 1.39 | 3.71
| 1.32 | 3.59
| 1.22 | 3.5
| 1.12 | 3.33
| 1 | 3.12
| 0.91 | 2.91
| 0.815 | 2.67
| 0.71 | 2.41
| 0.6 | 2.04
| 0.51 | 1.64
| 0.4 | 1.09
| 0.3 | 0.7
| 0.2 | 0.427
| 0.152 | 0.285
| 0.127 | 0.23
| 0.105 | 0.22
| 0.09 | 0.17
| 0.08 | 0.148
| 0.07 | 0.114
| 0.065 | 0.106
| 0.06 | 0.103
| 0.055 | 0.1
| 0.05 | 0.09
| 0.046 | 0.086
| 0.04 | 0.076
| 0.036 | 0.066
| 0.03 | 0.049
| 0.023 | 0.038
| 0.02 | 0.033
| 0.016 | 0.03
| 0.01 | 0.016
| 0.005 | 0.004
| 0.002 | -0.006
| 0.007 | 0
4 | 0.01 | 0.003
| 0.015 | 0.01
| 0.02 | 0.024
| 0.0247 | 0.028
| 0.03 | 0.033
| 0.0356 | 0.044
| 0.04 | 0.05
| 0.045 | 0.064
| 0.051 | 0.075
| 0.056 | 0.085
| 0.06 | 0.09
| 0.0035 | -0.003
| 0.01 | 0.004
| 0.02 | 0.024
| 0.03 | 0.033
| 0.04 | 0.054
| 0.05 | 0.057
| 0.06 | 0.08
| 0.07 | 0.1
| 0.08 | 0.104
| 0.09 | 0.107
| 0.1 | 0.131
| 0.125 | 0.21
| 0.15 | 0.227
| 0.175 | 0.248
| 0.2 | 0.3
| 0.226 | 0.343
| 0.25 | 0.401
| 0.275 | 0.434
| 0.3 | 0.503
| 0.324 | 0.56
| 0.352 | 0.613
| 0.376 | 0.668
| 0.4 | 0.743
| 0.452 | 0.89
| 0.501 | 1.01
| 0.547 | 1.08
| 0.6 | 1.25
| 0.65 | 1.31
| 0.7 | 1.53
| 0.75 | 1.65
| 0.8 | 1.816
| 0.85 | 1.98
| 0.9 | 2.11
| 0.95 | 2.3
| 1.01 | 2.46
| 1.11 | 2.71
| 1.24 | 3.24
| 1.33 | 3.45
| 1.4 | 3.58
| 1.52 | 3.74
| 1.61 | 3.86
| 1.49 | 3.74
| 1.41 | 3.7
| 1.3 | 3.5
| 1.22 | 3.46
| 1.09 | 3.26
| 1 | 3.08
| 0.95 | 2.93
| 0.9 | 2.88
| 0.856 | 2.74
| 0.81 | 2.64
| 0.756 | 2.54
| 0.712 | 2.39
| 0.656 | 2.18
| 0.6 | 2.01
| 0.55 | 1.84
| 0.5 | 1.61
| 0.452 | 1.25
| 0.4 | 1.06
| 0.352 | 0.853
| 0.3 | 0.68
| 0.277 | 0.56
| 0.25 | 0.53
| 0.225 | 0.438
| 0.2 | 0.387
| 0.176 | 0.308
| 0.15 | 0.255
| 0.126 | 0.205
| 0.1 | 0.172
| 0.091 | 0.15
| 0.08 | 0.118
| 0.07 | 0.077
| 0.06 | 0.068
| 0.05 | 0.063
| 0.04 | 0.031
| 0.03 | 0.007
| 0.02 | 0
| 0.01 | -0.027
5 | 0.01 | 0.008
| 0.02 | 0.022
| 0.03 | 0.028
| 0.04 | 0.05
| 0.05 | 0.063
| 0.06 | 0.076
| 0.07 | 0.107
| 0.08 | 0.129
| 0.09 | 0.139
| 0.104 | 0.162
| 0.125 | 0.188
| 0.154 | 0.213
| 0.175 | 0.272
| 0.2 | 0.325
| 0.226 | 0.37
| 0.251 | 0.422
| 0.278 | 0.49
| 0.31 | 0.532
| 0.328 | 0.57
| 0.352 | 0.64
| 0.377 | 0.73
| 0.403 | 0.805
| 0.453 | 0.947
| 0.502 | 1.04
| 0.55 | 1.12
| 0.59 | 1.27
| 0.65 | 1.43
| 0.7 | 1.59
| 0.75 | 1.71
| 0.8 | 1.9
| 0.87 | 2.09
| 0.99 | 2.48
| 1.1 | 2.74
| 1.22 | 3.23
| 1.33 | 3.45
| 1.45 | 3.63
| 1.58 | 3.81
| 1.44 | 3.73
| 1.36 | 3.62
| 1.32 | 3.55
| 1.2 | 3.44
| 1.11 | 3.27
| 1 | 3.08
| 0.91 | 2.89
| 0.8 | 2.63
| 0.72 | 2.46
| 0.66 | 2.26
| 0.6 | 2.02
| 0.57 | 1.86
| 0.52 | 1.77
| 0.446 | 1.27
| 0.4 | 1.13
| 0.355 | 0.91
| 0.29 | 0.69
| 0.258 | 0.576
| 0.21 | 0.435
| 0.15 | 0.309
| 0.103 | 0.225
| 0.91 | 0.175
| 0.082 | 0.16
| 0.07 | 0.13
| 0.061 | 0.113
| 0.05 | 0.104
| 0.038 | 0.065
| 0.031 | 0.04
| 0.02 | 0.023
| 0.011 | -0.01